## Requirements Matrix
Document Name | Reference (MDS-REQ1) | Class A | Class B | Class C | Class D | % Progress
Declaration of Conformity | Annex 4 | ✅ | ✅ | ✅ | ✅ | 30
Essential Principles Checklist | Annex 1 | ✅ | ✅ | ✅ | ✅ | 35
PMS Plan | Annex 6 | ✅ | ✅ | ✅ | ✅ | 40
Risk Management File | ISO 14971 / Annex 7 | ✅ | ✅ | ✅ | ✅ | 45
Clinical Evaluation | Annex 14 / MEDDEV 2.7.1 | ⚠️ | ✅ | ✅ | ✅ | 50
Labeling | Annex 3 | ✅ | ✅ | ✅ | ✅ | 55
UDI Submission | Annex 9 | ⚠️ | ✅ | ✅ | ✅ | 60

## Gap Analysis
Item | Observation | Regulatory Impact
IFU | Missing IFU contradicts Annex 1 Clause 23.4 | Registration not allowed
Label | No labeling format provided | Blocks UDI registration and delays submission
UDI-DI | UDI codes not assigned to models | GHAD submission not accepted
Biocompatibility Testing | No evidence of material safety | Direct violation of Annex 1 Clause 10
Label Validation Report | Functional verification not provided | Risk of non-compliance with ISO 15223 and 20417
Manufacturer License | Not attached | Required as legal proof per Annex 3

## Client Submission Tracker
Document | Submitted | Regulatory Technical Review
Declaration of Conformity | ✅ Yes | ✔️ Accepted – Includes ISO 13485, ISO 14971, ISO 10993, and 20417. Missing authorized signatory and clear date.
Essential Principles Checklist | ✅ Yes | ✔️ Good – Linked to technical standards, but some fields marked 'NA' without justification.
PMS Plan | ✅ Yes | ✔️ Compliant with ISO 20416. Includes data sources and timeline, but no tracking for product variants.
BENEFIT File | ✅ Yes | ✔️ General benefit-risk justification – acceptable for simplified clinical assessment, not a replacement for Biocompatibility.
IFU (Instructions for Use) | ❌ No | ❌ Not submitted at all.
Label | ❌ No | ❌ Not provided – Arabic and English versions required.
Label Validation Report | ❌ No | ❌ Not available.
UDI List | ❌ No | ❌ UDI-DI list not submitted for each model.
Biocompatibility Reports | ❌ No | ❌ Not submitted – ISO 10993-5, -10, -11 test results missing.

## Dashboard Summary
Status | Count
Received | 4
Not Received | 5